Metadata: SBSG Summer 2008 Data.xls 


METADATA

SBSG Summer 2008 Dataset

Senator Beck Stream Gauge (SBSG)

Senator Beck Basin Study Area (SBBSA), San Juan Mountains 

I. Senator Beck Stream Gauge (SBSG)

1) Site

a) Senator Beck Stream Gauge (SBSG)

i) RF Station ID = 4

(1) 1 hour data arrays:  Flag field = 401

(2) 24 hour data arrays: Flag field = 424

b) Site unlocked; datalogger enclosure locked with keyed padlock

c) Lat/Lon: 37( 54' 24" N, 107( 42' 31" W; NAD 1927

d) Elevation: 11,020 feet; 3,359 meters

e) Aspect: not applicable

f) Slope Angle: not applicable

g) Photographs: http://www.snowstudies.org/senbeck_sa.html
h) Changes to site: flume tailrace completed in summer 2005 

i) Comments: The broadcrested weir is equipped with a staff gauge for stage height and with automated sensors for stage, water temperature, and water electrical conductivity. 

2) Operation Status

a) Weir calibrated during summer of 2004 and early spring 2005 using manual discharge measurements performed at the weir at various stage heights by Colorado Division of Water Resource.  Those data were used by the weir contractor, R&M Construction, to conform the "Q" function for the actual construction and a discharge table for the weir was finalized using the equation:


[image: image1.wmf](

)

112

u

QKhK

=+


where:
h1 = stage height

K1 = 11


K2 = .07


u  = 2.75

b) Automated sensors operated for collection of 'Summer' datasets only, (spring through late fall), due to ice damage potential. 

c) Summer 2008 dataset from 1000 hrs March 15, 2008 (DOY 75) through 2400 hrs November 23, 2008 (DOY 328).
d) Significant stream bed movement during spring runoff required removing materials and restoring the gauge pool to its nominal depth on June 27, 2008.
e) Water temperature and EC data briefly interrupted for one hour on September 6, 2008 (DOY 250) during servicing.
f) All data collected on Mountain Standard Time, by day-of-year (DOY)

i) System clocks not switched from/to Mountain Daylight Time in Fall or Spring

ii) 24 hour summary datasets (Flag = 424) are calendar day
3) Type of Site

a) Study Plot: SBSG is located at the Basin pourpoint in a small gorge cut through a bedrock outcrop.

4) Power

a) Solar powered by 10 watt photovoltaic panel providing regulated charge to 8 amp hour lead-acid 12V DC battery through September 6, 2008; second 10 watt photovoltaic panel added to power supply on September 6.  Larger 12 amp-hour battery pack installed at shut-down of system on November 23, 2008.
5) Sensors

a) Stage Height 

i) Properties

(a) Make: Druck (GE Infrastructure Sensors)

(b) Model: PDCR 1830-8388 transducer (CS420) 

(c) Serial Number: 1892202

(d) Type: 2.5 psi transducer

ii) Installation

(a) Located two feet upstream of weir crest (centerline) and 0.1 feet below weir notch base, in open-top PVC pipe 'well' covered by nylon mesh screen

(b) Service/calibration dates: new instrument at installation in summer 2004; function tested OK by GE April 21, 2005. 

(c) Comments: sensor cable housed in flexible PVC and steel conduit to desiccant enclosure on mast.

b) Water Temperature and Electrical Conductivity

i) Properties

(a) Make: Campbell Scientific

(b) Model: CS547A

(c) Serial Number: 3997

(d) Type: 

ii) Installation

(a) Located in perforated PVC mounted horizontally on weir's upstream face below notch level

(b) Date installed: operated during summer of 2004

(c) New sensor when installed in spring 2004 

(d) Comments: EC sensor was scrubbed at 1030 hrs on September 6, 2008 (DOY 250) and showed a subsequent 0.003 mS/cm drop in readings
6) Datalogger

a) Make: Campbell Scientific

b) Model: CR10X

c) Serial Number: 39210

d) Type: fully programmable measurement and control system with ring memory

e) Acquisition date: May 2004

f) Date installed at SBSG: Spring 2004

g) Service/calibration dates: new when installed in May, 2004

h) Comments: extended temperature tested (-55 to +85C)

7) Data Retrieval

a) Radio telemetry using phone-to-RF base station; no repeater utilized

i) Campbell Scientific model RF 310M modem; serial #2532

ii) Model RF310 Maxon SD-125 V2 VHF radio: serial #040106488

iii) Antennex Yagi type antenna with 9' coax cable, two Type N connectors; antenna serial #38089

iv) Poly-Phaser IS-50NX-C2 in-line antenna surge protector 

8) Software

a) Campbell Scientific LoggerNet 2.1c

b) One SBSG Edlog (LoggerNet 2.1c) program versions (SBSG08V1) utilized during Summer 2008. 

9) Observer Contact Information

a) Name:  Chris Landry - Executive Director

b) Agency:  Center for Snow and Avalanche Studies

c) Address:  PO Box 190, Silverton, CO, USA 81433

d) Telephone:  (970) 387-5080

e) Email:  clandry@snowstudies.org
f) Website: http://www.snowstudies.org/index.html



(continued)

10) Data processing and output methods (data points as labeled in SBSG Summer 2008.xls Excel Spreadsheet):

a) Measurement execution interval:  5 seconds

b) Measurements: see table below for location labels per program SBSG08V1.CSI 
	Spreadsheet Labels 

(some abbreviations expanded)
	Loc'n

#
	LoggerNet 2.1c

Ouput Instruction
	Type of

Measurement
	Notes

	Flag
	1
	P80
	na
	400 = 10 minute array

401 = 1 hour array

424 = 24 hour array

	Year
	2
	P77
	Na
	Calendar year 

	Day of Year
	3
	P77
	na
	Sequential day of calendar year

	Time MST
	4
	P77
	na
	Mountain Standard Time at end of array period, in military time; previous day at midnight, 2400 at midnight

	EC (no T correction)
	5
	P70
	Sample
	EC uncorrected for water temperature; units are mScm-1

	Water Temp C
	6
	P71
	Average
	Array average water temperature (C)

	EC Correct 25C
	7
	P70
	Sample
	EC corrected to 25 deg C; units are mScm-1

	Water Depth ft
	8
	P71
	Average
	Stage height in feet

	Calculated “Q” Discharge CFS
	9
	Na
	Calculated field
	Field is calculated in sequence of instructions using derived formula shown in Part 2)a of Metadata

	System Voltage
	10
	P70
	Sample
	Voltage at datalogger


4 of 4

_1192693847.unknown

